Date: 20.10.1992

Version: SOMMER 1990, SOMMER 1992

Last modification: 02 June 1993

Author: Thomas Wieder

E-mail: WIEDER@PHYSIK.UNI-KASSEL.DE

0. Short Description

SBGBBG evaluates the full strain or stress tensor of a polycrystalline

material from X-ray reflection positions.  To obtain the strain (stress)

tensor, the user has to provide the reflection positions obtained from

an X-ray diffraction measurement, according to the well known

PHI-PSI-procedure.  Furthermore, the user has to provide either the

strain free (ideal) reflection positions or the strain free (ideal)

lattice constants, together with the crystal system and the reflection

indices.  The accuracy of the result depends strongly on the accuracy of

the strain free values.  To obtain the stress tensor, the user also has

to provide the X-ray elastic constants 1/2S2(HKL) and S1(HKL).

1. The SBGBBG Process

The purpose of SBGBBG is to evaluate the full strain tensor or

alternatively the full stress tensor of a polycrystalline sample from

X-ray diffraction data.  The data may be sampled with a diffractometer

in Seemann-Bohlin geometry (SBG) or with a diffractometer in

Bragg-Brentano geometry (BBG).  Therefore the name of the program is

composed from SBG and BBG, resulting in SBGBBG.

Strains E(HKL) are found from the change in lattice planes distances

from their unstrained value D0(HKL) to D(HKL) according to

E(HKL)=(D(HKL)-D0(HKL))/D0(HKL)                                    (1)

where HKL are the lattice planes indices.

The value of D0 must be known!  Thus, one must have an unstrained

sample, which is often not the case.  The accuracy of the results of

SBGBBG depends strongly on the correctness of the value for D0.  Under

certain circumstances it is possible to find D0 even from the strained

sample, but SBGBBG does not support this procedure.  Consult the

literature for this problem.

The value for D(HKL) is obtained by the X-ray measurement on the

strained sample.  To find the full strain tensor <E(I,J)> one must carry

out measurements in different directions with respect to the sample

coordinate system.  The azimuth angle PHI and the tilting angle PSI

indicate these directions.  In fact, one measures D(HKL,PHI,PSI) and

yields E(HKL,PHI,PSI) as measured strain.

The measurement is done in the following two ways, depending on

the geometry used.

For SBG: First one sets an incidence angle GAM, which remains unchanged

during the entire measurement.  For example one fixes GAM at GAM=6 deg.

Then the sample is placed into the sample holder with an azimuthal

orientation, which may be given arbitrarily the value PHI=0 deg.  Then a

scan over the accessible THETA range is made and the positions

THETA(HKL) of the occuring (HKL)-reflections are recorded.  According to

the SBG, each reflection occurs under its own tilting angle PSI(HKL).

For SBG one does not need to tilt the sample.  After the scan one

rotates the sample in PHI-direction to another PHI angle, for example

PHI=45 deg.  Then one does a new scan.  In this manner one proceeds for

NSCAN scans.  It is recommended to do at least NSCAN=7 scans.

For BBG: At first the sample is placed into the sample holder with an

azimuthal orientation, which may be given arbitrarily the value

PHI=0 deg.  Then the sample is tilted by a tilt angle PSI.  Then

a scan over the accessible THETA range is made and the positions

THETA(HKL) of the occuring (HKL)-reflections are recorded.  After the

scan one chooses another tilting angle PSI and does a new scan.  It is

recommended to vary also the azimuthal angle PHI, by rotating the sample

in PHI-direction to another PHI angle, for example PHI=45 deg.  Then one

does a new scan.  In this manner one proceeds for NSCAN scans.  It is

recommended to do at least NSCAN=6 scans.

To obtain the strain tensor {E(I,J)} (where the {} indicate the tensor

character), one has to transform the measured strain values

E(HKL,PHI,PSI) into the averaged components <E(I,J)> (where the <>

indicate the averaging) of {E(I,J)}. The transformation equation is:

E(HKL,PHI,PSI) =

<E(1,1)> COS(PHI) SIN(PSI) + <E(1,2)> SIN(2PHI) SIN(PSI) +

<E(1,3)> COS(PHI) SIN(2PSI) + <E(2,2)> SIN(PHI) SIN(PSI) +

<E(2,3)> SIN(PHI) SIN(2PSI) + <E(3,3) COS(PSI)**2                   (2)

Each measured E(HKL,PHI,PSI) provides one equation of the form (2).

Thus one has to solve an equation system for the determination of the 6

tensor components <E(I,J)>.  This is what SBGBBG mainly does.

Since one has to determine 6 components one needs at least 6 independent

equations, in other words 6 independent measurements.  One has to

measure at least 6 different D(HKL,PHI,PSI) which is easily achieved by

varying the parameters HKL and/or PHI and/or PSI, as described.  For SBG

we recommend to do at least 7 measurements, because for SBG the program

will also calculate the unknown radial sample displacement (the sample

never exactly touches the focusing circle of the diffractometer).

The <E(I,J)> are averages, because the dependency on (HKL) does not

appear on the right hand side of (2).  The averaging is based on the

mechanical anisotropy of the material.  If a material is mechanically

anisotropic, then it will have different strains in different lattice

directions [HKL].  Thus the measured strain depends on (HKL),

as it is written here.  If the material under investigation is

mechanically isotropic, then one may drop the dependency on (HKL) and

the average will be the exact solution.

In the usual stress measurement by use of BBG, one utilizes only one

(HKL)-reflection and does not calculate the strain tensor, but directly

the stress tensor {S(I,J)}. The anisotropy is taken into account by the

X-ray elastic constants 1/2S2(HKL) and S1(HKL), which are used to

convert the measured strains into stresses.

Thus by use of X-ray elastic constants, we can calculate directly the

stress tensor and can avoid the averaging about the anisotropy.  The

corresponding transformation equation form measured E(HKL,PHI,PSI) to

the stress tensor components S(I,J) is found from (2) by use of Hooke's

law and is

E(HKL,PHI,PSI) =

S(1,1) ( 1/2S2(HKL) COS(PHI) SIN(PSI) + S1(HKL) ) +

S(1,2) 1/2S2(HKL) SIN(2PHI) SIN(PSI) +

S(1,3) 1/2S2(HKL) COS(PHI) SIN(2PSI) +

S(2,2) ( 1/2S2(HKL) SIN(PHI) SIN(PSI) + S1(HKL) ) +

S(2,3) 1/2S2(HKL) SIN(PHI) SIN(2PSI) +

S(3,3) ( 1/2S2(HKL) COS(PSI)**2 + S1(HKL))                           (3)

Note that the (HKL) dependency appears on both sides, but here one has

to know 1/2S2(HKL) and S1(HKL) for each (HKL)!  The X-ray elastic

constants can be either measured (best way) or calculated.  Two

considerations enter into the calculation.

In the case of a single crystal one has for the crystallographic

direction [hkl] well defined single crystal elastic constants, the so

called compliance moduli SM(N,M,O,P) (Do not bother about the indices!).

Now consider a vector V, which moves along a certain direction within a

polycrystalline sample.  The vector will cross several crystallites,

which will have different crystallographic orientations with respect to

V.  The mechanical behavior of the polycrystalline sample along this

direction (and along all others) will therefore somehow depend on the

average of the compliance moduli.  Thus, one has to calculate volume

averages of the compliances.  There are two approximations to do that,

namely the Voigt average and the Reuss average.  The Voigt calculation

assumes constant strain within all grains regardless of their

orientation.  The Reuss calculation assumes constant stress along a

certain direction.  The truth is somewhere in between, as the experiment

shows.

The second consideration for the calculation of the X-ray elastic

constants takes the orientation of the reflecting crystallites into

account.  Not all crystallites contribute to the reflection, but only

those in reflection position.  This affects the volume average

of the elasticity.  One averages not over the entire volume, but only

over the volume of reflecting crystallites.  Note that the orientation

of reflecting crystallites within the sample depends on the

diffractometer geometry used.  Thus, X-ray elastic constants for SBG and

BBG need not be the same for same material and reflections!  On the

other hand, the differences may be small.  If one uses the Voigt

approximation, then no differences in x-ray elastic constants between

SBG and BBG will arise.

Refer to I.C.  Noyan and J.B.  Cohen, Residual Stress, New York

(1987), Chapter 3 to learn how to calculate the 1/2S2(HKL) and S1(HKL)

from the elastic compliances.  SBGBBG always asks for the

1/2S2(HKL) and S1(HKL).  If they are unknown, just provide an isotropic

(averaged) value, but then you can not calculate the stress tensor for

mechanically anisotropic material.  The following equations may be

useful for that:

1           1 + NUE                          NUE

- S2(HKL) = -------(HKL)   and  S1(HKL) = - -----(HKL)             (4)

2             EM                              EM

where NUE is Poisson's ration and EM is Young's modulus. For isotropical

material NUE and EM do not depend on (HKL).

To calculate the strain tensor or stress tensor choose the option TENSOR

in the menu EVALUATION.  The results are printed into a file named

PHIPSI.OUT.  Rename the file or print it if you want to keep the

results, since SBGBBG will overwrite PHIPSI.OUT the next time.

The E(HKL,PHI,PSI) are calculated from the measured reflection positions

THETA(HKL,PHI,PSI).  Accordingly one has to input this

THETA(HKL,PHI,PSI) into SBGBBG.  Some additional data are also

necessary.  SBGBBG asks for them if one does interactive input via the

terminal.  This is the best for the first time one uses SBGBBG.  Also it

is somewhat tedious.  For interactive input choose option TERMINAL in

the menu DATA.  It is also possible to input data from a file.  The data

has to be written in a specific order into the file.  SBGBBG does this

once you have given to it data interactively.  So you can produce a

proper sample file and edit it with an Ascii editor of your choice to

input different data into SBGBBG.

Some further remarks about the strategy of data input are in order.

Most of the input is checked for reasonable values.  That is, if you

input physical values, they will be checked for lying in a reasonable

range.  If data are recognized to be senseless, then SBGBBG will prompt

for new ones.  Generally SBGBBG will require you to answer its

questions.  In cases where you do not know the exact values required,

just input some reasonable one.  Be aware, that as a result some

subroutines will not give meaningful results.  If you are sick of

answering SBGBBG, then you eventually have to leave the program

forcefully by typing the escape-key or something similar.

If you mistyped during interactive input, then do not stop but continue.

After the complete input, you can correct the data through the EDITOR.

The rather primitive EDITOR shows all given data one after another on

the terminal and asks you, whether the value should be corrected or not.

1.2 Some more theory

1.2.1 FCC Lattices

For face-centered cubic materials (fcc) the stacking faults contribute

to the reflection shift.  Three types of stacking faults are relevant:

(a) intrinsic faults, also called deformation or single stacking faults,

(b) extrinsic or double stacking faults, and (c) twin stacking fault.

Intrinsic and extrinsic stacking faults contribute to peak shift

according to the theory.  In addition, intrinsic faults cause

symmetrical reflection broadening, extrinsic and twin fault provoke

asymmetrical broadening.  The peak shift due to extrinsic faults has

been verified experimentally in a number of pure metals (Cu, Ag, Au, Pb,

Ni) and alloy systems.  The theoretical consideration of the influence

of twin faults and extrinsic faults indicate that the corresponding

X-ray line broadening should be asymmetric in both cases and that the

peak shifts resulting from the presence of extrinsic faults should occur

in a direction opposite to the shift produced by intrinsic faulting.

SBGBBG is capable to take into account peak shift due to stacking

faults for fcc-materials.  To tell SBGBBG whether a cubic lattice is

fcc or not a control variable named KFZ must be set (5th line of input

file).  One has to type 1 for a fcc lattice and 0 or something else

for a non-fcc lattice.  So in case of a cubic fcc lattice the

corresponding input line would be

1 1                              lattice type: 1=cubic, 2=hexagonal, KFZ

For a cubic bcc lattice it would be

1 0                              lattice type: 1=cubic, 2=hexagonal, KFZ

For any other lattice no second number is necessary. If one wants to omit

account for fcc even in case of a fcc lattice, then one just types 0 as centering

number.  This is recommended for low accuracy measurements.

1.2.2 Gradients

There has been argumentation about the possibility of strain/stress gradients.

Mostly gradients orthogonal to the surface along the depth z are meant, but

also gradients along the surface can be thought of.  Brakman

(C.M. Brakman, J. Appl. Cryst. vol. 16 (1983), p 325-340) explains that

only a gradient in S(33) may be existent.  On the other hand, Wern (H. Wern,

Strain, November 1991, p.  127-136) allows for a variation of all

strain/stress tensor components along all three sample dimensions x,y,z.

SBGBBG owns an additional feature to calculate strain/stress gradients along z.

For that purpose the strain tensor is developed into a Taylor series with

respect to z.  Only the first two powers E(1,1)(Z=0), E(1,2)(Z=0),

E(1,3)(Z=0), E(2,2)(Z=0), E(2,3)(Z=0), E(3,3)(Z=0) and dE(1,1)/dZ,

dE(1,2)/dZ, dE(1,3)/dZ, dE(2,2)/dZ, dE(2,3)/dZ, dE(3,3)/dZ are taken

into consideration.  A straight line describes the gradient.  The

E(I,J)(Z=0) give the strain values at the surface, note that for the

strain the (3,3)-component may be unequal to zero in contrast to the

stress.  The dE(I,J)/dZ give the slopes of the lines.  They have units

of reciprocal length.  The user has to provide SBGBBG with the linear

absorption coefficient MUE in units of 1/cm (as MUE is usually

tabulated).  SBGBBG converts MUE into 1/microns (1 micron = 1.0E-06 m).

So the dE(I,J)/dZ are in 1/microns.

Up to now, only the gradient can be calculated for SBG.  The

calculation needs measurements made at different depthes, which can be

accomplished by varying the incidence angle GAM.  You have to carry out

scans made at several values of GAM and you have to type these scans have

into one input file.  SBGBBG asks you in case of SBG whether you would

like to calculate the gradient.  For more details see SUBROUTINE

GRADIENT.

In the current version, no correction for radial displacement is made

during the calculation of a depth-dependent tensor!

1.2.3 Radial Displacement

In the case of SBG, the program also gives a calculated value for the

most serious experimental error, namely the radial sample

displacement.  The tensor calculation corrects for this error, which

guaranties an improved accuracy of the results.  To obtain correction

for radial displacement the flag ICORR must be set equal to 1.  To

exclude this correction ICORR must be set equal to 0.  See sample input

files, fourth line, last number.

1.3 Examples

Two examples are provided in the following, one for SBG and one for BBG.

The input file is in free format.  It may have any name.  SBGBBG asks

for the file name.  Comments are written in lower capitals, they usually

do not appear in the files.  You may leave the comments in the input

file as an aid.

Give all angles in degrees.  The measured Bragg angles THETA and the

ideal Bragg angles THETAI are on a THETA-scale, not on a 2THETA-scale.

Please pay attention in case of BBG, where one usually inputs according

to a 2THETA-scale.  Most lengths in nm, but sometimes in mm or in 1/cm,

whereby this is indicated in the comments.  Units are given in square

brackets ([] means the quantity has no units, [a.u.] means arbitrary

units).

The results of the tensor calculation are printed into the file

PHIPSI.OUT.

We give some explanations of the printed results here.  The least

squares residual vector tells something about the quality of the

solution.  Its components should be as small as possible.  If they are

greater than the components of the non-diagonalized strain (stress)

tensor, then the results are meaningless, probably the measurement was

not accurate enough or the input was incorrect.  Usually they will have

values around 1.0E-03 (for strains) or less.  After diagonalization the

non-diagonal tensor elements (E(1,2) or S(1,2), E(1,3) or S(1,3), E(2,3)

or S(2,3)) should be almost equal to zero, values in the order of

1.0E-06 (for strains) or less are acceptable.  Of interest are the

diagonal elements of the diagonalized tensor, since these are the

principal axes values.

What you at least have to look for are E(1,1) (or S(1,1)), E(2,2)

(or S(2,2)) and E(3,3) (or S(3,3)) after diagonalization.  For the

strain tensor, the components should be in the order of 1.0E-02 or less

and may be positive or negative.  If their absolute values are much

greater than lets say 5.0E-02, then something went wrong.  In most

cases this means that one used wrong values for the ideal (strain

free) lattice constants or the ideal (strain free) Bragg angles, or

the measurement was not good enough.  For the stress tensor, the

stresses should not exceed the yield strength of the material.  Since

above the yield strength considerable plastic deformation begins, the

material can not support elastically higher stresses, but SBGBBG

assumes exclusively elastic deformation.  On the other hand, thin

films may exhibit considerable higher stresses than bulk solids of the

same material.

SBGBBG also calculates the eigenvectors of the tensor.  Their components

are the cosines of the angles between corresponding axes of the tensor

coordinate system and the sample coordinate system.  If the first

element of the first vector, the second element of the second vector and

the third element of the third vector are (almost) equal to one and all

other elements are (almost) equal to zero, then the principal axes of

the tensor are parallel to the sample coordinate system.

1.3.1 An example for SBG

Here the input file:                      Comments:

JULY 1992                                 date

P1E3  (TiN)                               sample name

0.1540510        788.4                    LAMBDA [nm], MUE [1/cm]

1     57.3   1                  geometry: 1=SBG 2=BBG, RADIUS [mm], ICORR

1            0              lattice type: 1=cubic, 2=hexagonal, KFZ

0.422220       0.422220       0.422220    lattice constants a, b, c [nm]

90.00000       90.00000       90.00000    cell angles alpha, beta, gamma

6                                     total number of scans NSCAN

3.0     .00000                            incident angle GAM, PHI

0      0.0                                just put two zeros here

5                                     number of reflections in the scan

21.38800       18.38800       21.39900    THETA,  PSI,  strain-free THETA0

2          0          0             H, K, L

.26300     987.00                    FWHM [deg.], PINT [a.u.]

0.0000048      -0.0000008             1/2S2, S1 elastic constants [mm**2/N]

31.06800       28.06800       31.06400    THETA,  PSI,  strain-free THETA0

2          2          0             H, K, L

.22900     517.00                    FWHM [deg.], PINT [a.u.]

0.0000048      -0.0000008             1/2S2, S1 elastic constants [mm**2/N]

37.18700       34.18700       37.23300

3          1           1

.35100    147.00

0.0000048      -0.0000008

54.52400       51.52400      54.67300

4          2           0

.46600     119.00

0.0000048      -0.0000008

63.12200       60.12100      63.34600

4          2          2

.62100     125.0

0.0000048      -0.0000008

3.0   45.00000                            here the next scan (at PHI=45) begins

0      0.0                                just put two zeros here

5                                     number of reflections in the scan

21.38900       18.38900       21.39900    THETA,  PSI,  strain-free THETA0

2          0          0             H, K, L

.25500     114.00                    FWHM [deg.], PINT [a.u.]

0.0000048      -0.0000008             1/2S2, S1 elastic constants [mm**2/N]

31.07400       28.07400       31.06400    THETA,  PSI,  strain-free THETA0

2          2          0             H, K, L

.22900     423.00                    FWHM [deg.], PINT [a.u.]

0.0000048      -0.0000008             1/2S2, S1 elastic constants [mm**2/N]

37.18600       34.18600       37.23300

3          1           1

.35100    146.00

0.0000048      -0.0000008

54.51600       51.51600      54.67300

4          2           0

.48300     135.0

0.0000048      -0.0000008

63.13200       60.13200      63.34600

4          2          2

.22800     167.0

0.0000048      -0.0000008

3.0   90.00000                            here the next scan (at PHI=90) begins

0      0.0                                just put two zeros here

5                                     number of reflections in the scan

21.35900       18.35900       21.39900    THETA,  PSI,  strain-free THETA0

2          0          0             H, K, L

.22700    1151.00                    FWHM [deg.], PINT [a.u.]

0.0000048      -0.0000008             1/2S2, S1 elastic constants [mm**2/N]

31.04500       28.04500       31.06400    THETA,  PSI,  strain-free THETA0

2          2          0             H, K, L

.22900     431.00                    FWHM [deg.], PINT [a.u.]

0.0000048      -0.0000008             1/2S2, S1 elastic constants [mm**2/N]

37.15700       34.15800       37.23300

3          1           1

.35100    143.00

0.0000048      -0.0000008

54.49000       51.49000      54.67300

4          2           0

.38600     151.0

0.0000048      -0.0000008

63.15500       60.15500      63.34600

4          2          2

.44700      87.00

0.0000048      -0.0000008

3.0   180.0000                            here the next scan (at PHI=180) begins

0      0.0                                just put two zeros here

5                                     number of reflections in the scan

21.38500       18.38500       21.39900    THETA,  PSI,  strain-free THETA0

2          0          0             H, K, L

.23300     690.00                    FWHM [deg.], PINT [a.u.]

0.0000048      -0.0000008             1/2S2, S1 elastic constants [mm**2/N]

31.05900       28.05900       31.06400    THETA,  PSI,  strain-free THETA0

2          2          0             H, K, L

.22800     446.00                    FWHM [deg.], PINT [a.u.]

0.0000048      -0.0000008             1/2S2, S1 elastic constants [mm**2/N]

37.17900       34.17900       37.23300

3          1           1

.33900    144.00

0.0000048      -0.0000008

54.51500       51.51000      54.67300

4          2           0

.55500     120.0

0.0000048      -0.0000008

63.09200       60.09200      63.34600

4          2          2

.62000      97.00

0.0000048      -0.0000008

3.0   225.0000                            here the next scan (at PHI=225) begins

0      0.0                                just put two zeros here

5                                     number of reflections in the scan

21.35300       18.35300       21.39900    THETA,  PSI,  strain-free THETA0

2          0          0             H, K, L

.22800     450.00                    FWHM [deg.], PINT [a.u.]

0.0000048      -0.0000008             1/2S2, S1 elastic constants [mm**2/N]

31.02200       28.02200       31.06400    THETA,  PSI,  strain-free THETA0

2          2          0             H, K, L

.22800     515.00                    FWHM [deg.], PINT [a.u.]

0.0000048      -0.0000008             1/2S2, S1 elastic constants [mm**2/N]

37.15400       34.15400       37.23300

3          1           1

.36300    184.00

0.0000048      -0.0000008

54.50600       51.50600      54.67300

4          2           0

.43900     139.0

0.0000048      -0.0000008

63.16400       60.16400      63.34600

4          2          2

.68600     116.00

0.0000048      -0.0000008             1/2S2, S1 elastic constants [mm**2/N]

3.0   270.0000                            here the next scan (at PHI=270) begins

0      0.0                                just put two zeros here

5                                     number of reflections in the scan

21.37100       18.37100       21.39900    THETA,  PSI,  strain-free THETA0

2          0          0             H, K, L

.24700     613.00                    FWHM [deg.], PINT [a.u.]

0.0000048      -0.0000008             1/2S2, S1 elastic constants [mm**2/N]

31.04900       28.04900       31.06400    THETA,  PSI,  strain-free THETA0

2          2          0             H, K, L

.27500     477.00                    FWHM [deg.], PINT [a.u.]

0.0000048      -0.0000008             1/2S2, S1 elastic constants [mm**2/N]

37.17700       34.17700       37.23300

3          1           1

.03190    148.00

0.0000048      -0.0000008

54.52800       51.52800      54.67300

4          2           0

.47500     112.0

0.0000048      -0.0000008

63.13000       60.13000      63.34600

4          2          2

.75200     120.00

0.0000048      -0.0000008                 last line of input file

Further comments on the input:

LAMBDA = wavelength

MUE = linear absorption coefficient, not in use at the moment

RADIUS = radius of the diffractometer, needed for correction of radial

sample displacement, give RADIUS in mm.

lattice type: 1=cubic, 2=hexagonal, 3=monoclinic, 4=tetragonal,

5=rhombic, 6=rhombohedral, 7=triclinic

lattice constants: Give the strain-free lattice constants!  If you do

not know them, then just give some reasonable values here.  This has two

consequences: You must provide the strain-free values THETA0(HKL) of the

reflections which you use!  Furthermore you must choose the option

CALCULATION WITH IDEAL ANGELS when you are in subroutine STRAIN.

THETA = Bragg angle of the reflection

THETA0 = strain free Bragg angle of the reflection.  Give THETA instead

of THETA0, if THETA0 is unknown.  This implies two consequences: You

must provide the strain-free lattice constants a,b,c!  Furthermore you

must choose the option CALCULATION WITH IDEAL LATTICE CONSTANTS when you

are in subroutine STRAIN.

FWHM = reflection's full width at half height on the THETA scale (not in

2THETA)

PINT = integrated intensity of the reflection

1/2S2 : The value 1/2 = 0.5 is usually included in tabulated values. So

do not multiply the tabulated values by 1/2.

S1 : The values for S1 are usually given as negative values, according

to the minus sign in (4b).  SBGBBG also expect negative values.

There exists an advanced method for correction of the strain values. For

its use one have to give certain extrapolated error-free strain values.

Since we seldom use this method, just put the two zeros in the line

after the line with GAM and PHI, as it is done in the examples.

Here the result file (PHIPSI.OUT):

JULY 1992                                 date     (first line of the output)

(TiN)                               sample name

STRAIN TENSOR:

NUMBER OF EQUATIONS NZ=  30

THIS FOLLOWING EQUATION SYSTEM HAS BEEN SOLVED.

USUALLY ONE HAS NOTHING TO DO WITH THIS SYSTEM, SKIP IT!

0.0995089  0.0000000  0.5986881  0.0000000  0.0000000  0.9004912  0.9032677

0.0004899

0.2213883  0.0000000  0.8303626  0.0000000  0.0000000  0.7786117  0.5701268

-0.0001160

0.3157268  0.0000000  0.9296093  0.0000000  0.0000000  0.6842732  0.4459324

0.0010565

0.6128836  0.0000000  0.9741813  0.0000000  0.0000000  0.3871165  0.2324577

0.0018432

0.7518266  0.0000000  0.8639058  0.0000000  0.0000000  0.2481733  0.1614068

0.0019624

0.0497597  0.0995193  0.4233562  0.0497597  0.4233562  0.9004806  0.9032638

0.0004454

0.1107376  0.2214752  0.5872375  0.1107376  0.5872375  0.7785249  0.5701185

-0.0002898

0.1578553  0.3157106  0.6573240  0.1578553  0.6573240  0.6842894  0.4459333

0.0010794

0.3063738  0.6127476  0.6888947  0.3063738  0.6888948  0.3872524  0.2324615

0.0019421

0.3759963  0.7519926  0.6107368  0.3759964  0.6107369  0.2480074  0.1614025

0.0018748

0.0000000  0.0000000  0.0000000  0.0992061  0.5978770  0.9007939  0.9033779

0.0017816

0.0000000  0.0000000  0.0000000  0.2210550  0.8299149  0.7789450  0.5701582

0.0005504

0.0000000  0.0000000  0.0000000  0.3152564  0.9292358  0.6847436  0.4459592

0.0017454

0.0000000 -0.0000001  0.0000000  0.6123055  0.9744485  0.3876945  0.2324736

0.0022637

0.0000000 -0.0000001  0.0000000  0.7523391  0.8633075  0.2476609  0.1613924

0.0016733

0.0994775  0.0000000 -0.5986043  0.0000000 -0.0000001  0.9005225  0.9032790

0.0006235

0.2212578  0.0000000 -0.8301874  0.0000000 -0.0000001  0.7787422  0.5701391

0.0001448

0.3155970  0.0000001 -0.9295064  0.0000000 -0.0000001  0.6844030  0.4459396

0.0012402

0.6126455  0.0000001 -0.9742915  0.0000000 -0.0000001  0.3873544  0.2324619

0.0019545

0.7513893  0.0000001 -0.8644153  0.0000000 -0.0000001  0.2486108  0.1614200

0.0022252

0.0495717  0.0991435 -0.4226442  0.0495717 -0.4226441  0.9008566  0.9034008

0.0020487

0.1103610  0.2207219 -0.5865216  0.1103610 -0.5865214  0.7792780  0.5701899

0.0012169

0.1575957  0.3151914 -0.6570325  0.1575957 -0.6570324  0.6848085  0.4459619

0.0018144

0.3062888  0.6125776 -0.6889504  0.3062887 -0.6889503  0.3874225  0.2324661

0.0020658

0.3762375  0.7524748 -0.6103384  0.3762373 -0.6103383  0.2475252  0.1613885

0.0015945

0.0000000  0.0000000  0.0000000  0.0993313 -0.5982127  0.9006687  0.9033322

0.0012471

0.0000000  0.0000000  0.0000000  0.2211130 -0.8299928  0.7788870  0.5701529

0.0004346

0.0000000  0.0000000  0.0000000  0.3155645 -0.9294806  0.6844355  0.4459414

0.0012862

0.0000000  0.0000000  0.0000000  0.6129516 -0.9741498  0.3870484  0.2324558

0.0017937

0.0000000  0.0000000  0.0000000  0.7519624 -0.8637475  0.2480376  0.1614034

0.0018923

IN CASE OF SBG: CORRECTION FOR SAMPLE DISPLACEMENT:

CALCULATED RADIAL SAMPLE DISPLACEMENT =      0.17229E-02mm

THE RESIDUAL VECTOR IS:

-0.33270E-03

-0.74462E-03

0.32698E-03

0.42055E-03

0.13223E-03

-0.43085E-03

STRAIN TENSOR COMPONENTS BEFORE DIAGONALIZATION:

E(1,1)=      0.26710E-02

E(1,2)=     -0.23481E-04

E(1,3)=     -0.25840E-03

E(2,2)=      0.28181E-02

E(2,3)=      0.77006E-05

E(3,3)=     -0.93802E-03

***!!!*** THE MAIN RESULT: ***!!!***

STRAIN TENSOR COMPONENTS AFTER DIAGONALIZATION:

E(1,1)=      0.26850E-02

E(1,2)=      0.16690E-08

E(1,3)=      0.00000E+00

E(2,2)=      0.28224E-02

E(2,3)=     -0.48851E-12

E(3,3)=     -0.95644E-03

THE EIGENVECTORS ARE:

1. VECTOR         2.VECTOR         3.VECTOR

0.98169607E+00  -0.17670727E+00   0.71044646E-01

0.17726636E+00   0.98416162E+00  -0.15930060E-02

-0.69637910E-01   0.14157671E-01   0.99747187E+00       (last line)

1.3.2 An example for BBG (The measured reflection positions are of low

accuracy. However, they serve as an example)

Here the input file:                      Comments:

08.04.1991                                date, (first line of input file)

Cha67                                    sample name

0.2290920         1386.00               LAMBDA [nm], MUE [1/cm]

2   100.   0                  geometry: 1=SBG 2=BBG, RADIUS [mm], ICORR

1   0                     lattice type: 1=cubic, 2=hexagonal, KFZ

0.361500       0.361500       0.361500 lattice constants a, b, c [nm]

90.00000       90.00000       90.00000 cell angles alpha, beta, gamma

3                                     total number of scans NSCAN

6.00000        0.00000                incident angle GAM, PHI

0              0.00000                just put two zeros here

7                                     number of reflections in the scan

33.33700        0.00000       33.2863  THETA,  PSI,  strain-free THETA0

1          1          1             H, K, L

0.19329     1190.00000                FWHM [deg.], PINT [a.u.]

0.000008732       -0.00000223             1/2S2(HKL), S1(HKL) [mm**2/N]

33.31070        5.00000       33.2863  THETA,  PSI,  strain-free THETA0

1          1          1             H, K, L

0.17571      990.00000                FWHM [deg.], PINT [a.u.]

0.000008732       -0.00000223             1/2S2(HKL), S1(HKL) [mm**2/N]

33.30930       10.00000       33.2863  here the next scan (at PSI=10) begins

1          1          1             input as above

0.20865     1137.00000

0.000008732       -0.00000223

33.28552       20.00000       33.2863

1          1          1

0.56200     1541.00000

0.000008732       -0.00000223

33.32028       -5.00000       33.2863

1          1          1

0.17387      710.00000

0.000008732       -0.00000223

33.31963      -10.00000       33.2863

1          1          1

0.19263      622.00000

0.000008732       -0.00000223

33.31949      -15.00000       33.2863

1          1          1

0.20764      486.00000

0.000008732       -0.00000223

6.00000       45.00000

0              0.00000

6

33.31063        0.00000       33.2863

1          1          1

0.16177      542.00000

0.000008732       -0.00000223

33.32956        5.00000       33.2863

1          1          1

0.18000      936.00000

0.000008732       -0.00000223

33.24025       20.00000       33.2863

1          1          1

0.20859      694.00000

0.000008732       -0.00000223

33.33265       -5.00000       33.2863

1          1          1

0.17200      685.00000

0.000008732       -0.00000223

33.33099      -10.00000       33.2863

1          1          1

0.16973      685.00000

0.000008732       -0.00000223

33.33358      -15.00000       33.2863

1          1          1

0.19979      418.00000

0.000008732       -0.00000223

6.00000       90.00000                the last scan

0              0.0

6

33.32465        0.00000       33.2863

1          1          1

0.16650      714.00000

0.000008732       -0.00000223

33.28149        5.00000       33.2863

1          1          1

0.17840      786.00000

0.000008732       -0.00000223

33.218966       20.00000       33.2863

1          1          1

0.22680      948.00000

0.000008732       -0.00000223

33.29053       -5.00000        33.2863

1          1          1

0.17158      629.00000

0.000008732       -0.00000223

33.29476      -10.00000       33.2863

1          1          1

0.19380      521.00000

0.000008732       -0.00000223

33.29509      -15.00000       33.2863

1          1          1

0.21650      386.00000

0.000008732       -0.00000223          (the last line of the input file)

For BBG please set incidence angle GAM simply to 6.0 deg, it will be

ignored.

Here the result file (PHIPSI.OUT):

08.04.1991                                (first line of output file)

Cha67                                    sample name

STRAIN TENSOR:

NUMBER OF EQUATIONS NZ=  19

THIS FOLLOWING EQUATION SYSTEM HAS BEEN SOLVED.

USUALLY ONE HAS NOTHING TO DO WITH THIS SYSTEM, SKIP IT!

0.0000000  0.0000000  0.0000000  0.0000000  0.0000000  1.0000000 -0.0013478

0.0075961  0.0000000  0.1736482  0.0000000  0.0000000  0.9924039 -0.0006486

0.0301537  0.0000000  0.3420201  0.0000000  0.0000000  0.9698462 -0.0006114

0.1169778  0.0000000  0.6427876  0.0000000  0.0000000  0.8830222  0.0000208

0.0075961  0.0000000 -0.1736482  0.0000000  0.0000000  0.9924039 -0.0009032

0.0301537  0.0000000 -0.3420201  0.0000000  0.0000000  0.9698462 -0.0008860

0.0669873  0.0000000 -0.5000000  0.0000000  0.0000000  0.9330127 -0.0008823

0.0000000  0.0000000  0.0000000  0.0000000  0.0000000  1.0000000 -0.0006468

0.0037981  0.0075961  0.1227878  0.0037981  0.1227878  0.9924039 -0.0011499

0.0584889  0.1169778  0.4545195  0.0584889  0.4545195  0.8830222  0.0012243

0.0037981  0.0075961 -0.1227878  0.0037981 -0.1227878  0.9924039 -0.0012321

0.0150768  0.0301537 -0.2418447  0.0150768 -0.2418448  0.9698462 -0.0011879

0.0334936  0.0669873 -0.3535534  0.0334937 -0.3535534  0.9330127 -0.0012569

0.0000000  0.0000000  0.0000000  0.0000000  0.0000000  1.0000000 -0.0010194

0.0000000  0.0000000  0.0000000  0.0075961  0.1736482  0.9924039  0.0001279

0.0000000  0.0000000  0.0000000  0.1169778  0.6427876  0.8830222  0.0017900

0.0000000  0.0000000  0.0000000  0.0075961 -0.1736482  0.9924039 -0.0001124

0.0000000  0.0000000  0.0000000  0.0301537 -0.3420201  0.9698462 -0.0002249

0.0000000  0.0000000  0.0000000  0.0669873 -0.5000000  0.9330127 -0.0002336

THE RESIDUAL VECTOR IS:

-0.53050E-03

0.40598E-06

-0.17382E-03

-0.67114E-04

0.34189E-04

0.11966E-03

STRAIN TENSOR COMPONENTS BEFORE DIAGONALIZATION:

E(1,1)=      0.23574E-02

E(1,2)=     -0.26287E-02

E(1,3)=      0.83053E-03

E(2,2)=      0.16975E-01

E(2,3)=      0.10506E-02

E(3,3)=     -0.81732E-03

***!!!*** THE MAIN RESULT: ***!!!***

STRAIN TENSOR COMPONENTS AFTER DIAGONALIZATION:

E(1,1)=      0.22217E-02

E(1,2)=      0.85692E-09

E(1,3)=      0.00000E+00

E(2,2)=      0.17477E-01

E(2,3)=     -0.23059E-06

E(3,3)=     -0.11836E-02

THE EIGENVECTORS ARE:

1. VECTOR         2.VECTOR         3.VECTOR

0.94037795E+00  -0.16848508E+00  -0.29546890E+00

0.14566799E+00   0.98449033E+00  -0.97773246E-01

0.30735961E+00   0.48903450E-01   0.95033598E+00     (last line)

1.4 An example course through SBGBBG to obtain the strain tensor

Start SBGBBG

Main menu: Choose task DATA

menu DATA: Choose task FILE

Give file name

Give "1" to see a protocol of the read data on the terminal

Main menu: Choose EVALUATION

menu EVALUATION: Choose TENSOR

menu TENSOR: Choose CALCULATION OF STRAIN OR STRESS TENSOR

menu DEHNUNG: Choose CALCULATION WITH STRAIN FREE LATTICE CONSTANTS

menu PHIPSI: Choose STRAIN TENSOR FOR ISOTROPIC MATERIAL

main menu: Choose PROGRAM EXIT

Print the file PHIPSI.OUT

2. Compiling SBGBBG

To compile SBGBBG just immediately try to compile it without any source

code modifications.  The compilation of SBGBBG should not provoke error

messages, but result in an executable file.

If you use a personal computer the source code in file SBGBBG.FOR may be

too long for your compiler. Than simply spilt SBGBBG.FOR into two files

of more or less equal size named SBGBBG1.FOR and SBGBBG2.FOR and compile

and link them.

Two problems could arise: (a) during linking the linker could report an

unresolved reference, (b) when running SBGBBG it could happen that the

file names uses by SBGBBG do not fit to the operating system of your

computer. In the following these problems are explained and the way to

remove them is given.

SBGBBG contains one operating system dependent subroutine.  The purpose

of this subroutine (named BEFEHL) is to allow a call of the operating

system's commands without leaving SBGBBG.  The correct system command

call is enabled by informing SBGBBG about the running operating system.

The Fortran OPEN command for file opening has been used in a way that it

should work on most common computers without need of change.

Nevertheless the file names are often incompatible between different

operating systems.  SBGBBG "knows" the appropriate file name conventions

according to several operating systems (VAX VMS, Norsk Data Sintran,

Microsoft DOS, IBM AIX (UNIX)).

To tell SBGBBG which operating system command call and which file name

convention is to be used, do the following before compiling: enable the

source code line

C      SYSTEM='SINTRAN'

C      SYSTEM='DOS'

C      SYSTEM='VAX'

C      SYSTEM='UNIX'

containing the appropriate system name at the beginning of SBGBBG's

source code by removing the "C" in the first column of the line (and let

the other ones be disabled by the "C" in their first columns).  For

example to enable UNIX it has to be

C      SYSTEM='SINTRAN'

C      SYSTEM='DOS'

C      SYSTEM='VAX'

SYSTEM='UNIX'

Actually one of these convention will be enabled by default and it

should be the correct one for your computer, if you have told me your

operating system.

3. What else can be done by SBGBBG

SBGBBG has also some subroutines to calculate some other structural

quantities.  The results of the now mentioned subroutines will be found

in files with names "SUBROUTINE.OUT", e.g.  FBA.OUT.

Very useful is subroutine FBA.  It extrapolates from one scan the error

free value A(THETA=90) of the lattice constant of cubic material.

However, A(THETA=90) may be affected by strain.  FBA ask you which of

the scans should be used for the extrapolation.  Subroutine HEXAGON does

this for hexagonal material.

Also useful is subroutine SCHERRER.  It calculates the mean crystallite

diameters according to Scherrer's equation.  Therefore SBGBBG asks for

the full widths at half height (FWHM) of the reflections.  If you are

not interested in crystallite diameters, just input a reasonable

constant value for all FWHMs.  SCHERRER asks you, which of the scans

should be used for the calculation.  Results are in file SCHER-2.OUT

The subroutines FBE and ABILDER are of interest only for experts.

4. What else you should know before you use SBGBBG

The program has been used already in several locations by students of

physics and materials science.  It has been found to be useful, since

the calculation of the strain/stress tensor by hand is tedious.  That is

the entire reason for existence of SBGBBG.  Use SBGBBG, if you find it

convenient for tensor calculation and do not expect more from it.

The input to SBGBBG has been reported to be tedious.  That's right, but

I can't help it.  We have to tell the measured reflection positions and

what else is necessary to SBGBBG in any way.  The interface to SBGBBG is

the input file.  THE USER IS FREE TO WRITE HIS/HER OWN INPUT ROUTINE,

WHICH SETS UP THE INPUT FILE MORE CONVENIENTLY!  In fact this has been

done at the University Kassel.  An entire system of software carries out

the measurement, performs profile fitting and writes the input file for

SBGBBG.  Unfortunately this system is accommodated to the specific

diffractometer there and is not suited for distribution.

If you intend to use SBGBBG, you should be familiar with strain/stress

measurement by X-ray diffraction!  Here some literature:

J.B. Cohen and H. Doelle: Stress Analysis from Powder Diffraction

Patterns; National Bureau of Standards Special Publication 567,

Proceedings of symposium on Accuracy in Powder Diffraction held at NBS,

Gaithersburg, USA, 1979, (1980), 453-477.

I. C. Noyan, J.B. Cohen: Residual Stress; Springer Verlag, Berlin

(1986).

R. Feder, B.S. Berry: Seemann-Bohlin X-ray Diffractometer for Thin

Films, J. Appl. Cryst. 3 (1970), 372-379.

E.L.  Haase, The determination of lattice parameters and strains in

stressed thin films using X-ray diffraction: Extensions, in: X-Ray and

Neutron Structure Analysis in Materials Science (Proc. Conf.), Karlovy

Vary, Tschechoslowakei, 5-9 October 1987, Plenum Press, New York,

(1989), 191-198.

A.J. Perry, V. Valvoda and D. Rafaja: X-ray residual stress

measurement in TiN, ZrN and HfN films using the Seemann-Bohlin method,

Thin Solid Films, 214 (1992), 169-174.

J. Zendehroud, T. Wieder, K. Thoma und H. Gaertner: Tiefenaufloesende

roentgenographische Dehnungsmessung an TiN-Schichten in Seemann-Bohlin-

Geometrie, Haertereitechnische Mitteilungen 48 (1991), 41-49.

This program description is not exhaustive, although SBGBBG is by no

means a huge or a complex program.  If you are trained in using programs

you will manage SBGBBG.

You will find additional explanations in the source text of SBGBBG or

even by using the program. Choose option INFO in the main menu.

4. Trouble with SBGBBG

4.1 Troubleshooting

Please forgive me, if you have bad experiences with SBGBBG!  Even

better, tell me what went wrong or what is annoying, so perhaps SBGBBG

can be improved.

Here my address for correspondence:

Dr. Thomas Wieder

University Kassel

Department of Physics

Experimentalphysik III

Wilhelmshher Allee 73

D-3500 Kassel

Germany

or use E-mail:

wieder@physik.uni-kassel.de

4.2 LEGAL MATTERS

I make no warranties, express or implied, that the program SBGBBG is

free of error, or is consistent with any particular standard of

merchantability, or that it will meet your requirements for any

particular application.  It (SBGBBG) should not be relied on for solving

a problem whose incorrect solution could result in injury to a person or

loss of property.  If you do use the program in such a manner, it is at

your risk.  The author disclaims all liability for direct or

consequential damages resulting from your use of the program.